Survey

This survey is used as a preliminary evaluation or requalification of your laboratory services.

Completion and submission of this survey does not constitute an approval of your company as an approved vendor.

Instructions

Please complete this survey in detail. Answers should reflect your current status.

Not all questions may apply to your situation; if not applicable, please complete with “NA”. Requested or additional documents may be added as an annex or submitted as separate documents. If additional information is provided to answer a question, please state in your response whether that information is annexed or sent as a separate document.

Please note that when comments are requested in the “please explain” section, an explanation is still required even if the answer is “No”.

This assessment will be used by Cerba Research for evaluation purposes only. All information will be treated as confidential.

Completed surveys can be emailed to: qavendormanagement@cerbaresearch.com

Thank you for taking the time to complete this questionnaire; please contact the Quality Department (qavendormanagement@cerbaresearch.com) if you require any further clarification.

Approval

As an authorized representative of my company, I hereby certify that this form is complete and accurate:

| Name: |  |
| --- | --- |
| Title: |  |
| Date: |  |
| Signature: |  |

| COMPANY PROFILE INFORMATION |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Company name | Shanghai Outdo Clinic Co., Ltd. |  |  |  |  |  |  |  |
| Website | N/A |  |  |  |  |  |  |  |
| Address | No. 918 Ziping Road, Zhoupu Town, Pudong New Area, Shanghai, China |  |  |  |  |  |  |  |
| QA main contact point | Linrong Liang |  |  |  |  |  |  |  |
| Function / job title | Quality Manager (Quality Responsible Person) |  |  |  |  |  |  |  |
| Phone / mobile No. | 13902559694 |  |  |  |  |  |  |  |
| E-mail | linrong_liang@shbiochip.com |  |  |  |  |  |  |  |
| Brief description of business | Shanghai Outdo Clinic Co., Ltd. is a wholly-owned subsidiary of Shanghai Outdo Biotech Co., Ltd., established in 2016 and located at Building 15, Lane 908, Ziping Road, Zhoupu Town, Pudong New Area, Shanghai. Registered capital is RMB 45.81 million. The site occupies 1,858.94 m², of which 892.29 m² is laboratory space. The company is an independent legal entity and a for-profit joint-stock medical institution. It performs medical laboratory and pathology testing in clinical microbiology, pathology and clinical cytomolecular diagnostics, and has established its quality management system in accordance with CNAS-CL02:2023 and related accreditation criteria. |  |  |  |  |  |  |  |
| How long has your company been providing services to the biotech/pharma industry supporting clinical trials? | Since 2019 |  |  |  |  |  |  |  |
| Company operational locations | ☐ Global ☐ Africa ☒ Asia-Pacific ☐ Central & South America ☐ Europe ☐ Middle East ☐ North America |  |  |  |  |  |  |  |
| Please provide areas of expertise | ☒ Special analysis ☒ Safety testing ☐ PK/PD ☒ Processing (aliquoting, PBMC isolation) ☒ Method validation Other, please specify: clinical pathology diagnostics (IHC & ISH); clinical bacterial culture (Helicobacter pylori culture and antimicrobial susceptibility testing); gut microbiome testing; cell culture and passaging; viral seed-stock preparation; viral titration; ELISA; TCID50 |  |  |  |  |  |  |  |
| Supplemental information relevant to services / testing provided to Cerba Research | Please provide information: RT-qPCR, sequencing, cell culture, turkey (avian) blood collection, virus culture, virus titration, TCID50, ELISA, EC50 |  |  |  |  |  |  |  |
| Turnover in the last 2 years | RMB 80 million |  |  |  |  |  |  |  |
| Number of employees | Quality staff |  | Laboratory staff |  |  | Administration / other |  | IT staff |
|  | 4 |  | 30 |  |  | 9 |  | 2 |
| Is your company a subsidiary of a parent company? |  |  |  |  | Yes No Yes No |  |  |  |
| If Yes, please state the name and address of the parent organization. |  |  |  |  |  |  |  |  |
| Shanghai Outdo Biotech Co., Ltd. No. 151 Libing Road, Zhangjiang, Pudong New Area, Shanghai, China Shanghai Outdo Biotech Co., Ltd. was established in December 2003 and is part of the Shanghai Biochip Co., Ltd. group. It is responsible for operation of the central platform, including the clinical-translation CBDTM paradigm. It is a national-level Specialized, Refined, Distinctive and Innovative “Little Giant” enterprise and a Shanghai key high-tech enterprise, operating under a dual-driver model with Outdo and the biochip platform. Shanghai Biochip Co., Ltd. was established in August 2001 by Shanghai Science and Technology Venture Capital (Group) Co., Ltd., Shanghai Zhangjiang (Group) Co., Ltd., relevant CAS Shanghai institutes, Shanghai Jiao Tong University and other universities, research institutes, hospitals and enterprises in Shanghai. With NDRC approval it constructs and operates the National Engineering Research Center for Biochip at Shanghai, one of the largest national engineering research centers in China’s biopharmaceutical field. The Center serves as the secretariat and chair organization of the National Technical Committee on Biobanking Standardization (SAC/TC559), the Biobank Branch of the China Medicinal Biotechnology Association (BBCMBA), and the Tumor Biospecimen Integration Research Branch of the China Anti-Cancer Association. |  |  |  |  |  |  |  |  |
| List main customers to which you have provided services in the past two (2) years, preferably in the same area/product range as the current request. |  |  |  |  |  |  |  |  |
| Company name: |  |  |  | Contact name: |  |  |  |  |
| N/A |  |  |  | N/A |  |  |  |  |
| Have you been inspected by the US FDA or other national regulatory authority? Please provide history of regulatory inspections with associated dates and results/reports. |  |  |  | Yes No Yes No If Yes, please provide the details below |  |  |  |  |
| Authority |  | Inspection date |  | Overall results / reports / outcome |  |  |  |  |
| NMPA |  | 2025.11.14-11.16 |  | Pass; the final result/report is retained by the Sponsor and Cerba |  |  |  |  |
| US FDA debarment list (Definition: list of persons who are debarred, meaning they are effectively forbidden from working in the drug industry) |  | Could you confirm that none of your employees is on the US FDA debarment list? |  | Yes No Yes No If No, please provide the details below |  |  |  |  |
| Is your company certified or accredited (e.g. ISO 15189, CAP or other in line with your services)? |  |  |  | Yes No Yes No ISO15189 |  |  |  |  |
| Standard |  | Certification / Accreditation body |  | Certified / Accredited since? |  |  | Expiration date |  |
| Please provide the certificate(s) |  |  |  |  |  |  |  |  |
| ISO15189 |  | China National Accreditation Service for Conformity Assessment (CNAS) |  | 2026-03-12 |  |  | 2030-01-10 |  |
| Overview of scope of the certification |  |  |  |  |  |  |  |  |
| Products / services |  |  |  |  |  |  | Within scope of certification |  |

| QUALITY SYSTEMS AND PROCEDURES |  |  |  |  |
| --- | --- | --- | --- | --- |
| ### Question | ### Yes | ### No | ### N/A | ### Comments / specification |
| ### Do you have a quality management system? |  |  |  | ### ISO15189 |
| ### If yes, please summarize your QMS framework. | The QMS framework is established in accordance with CNAS-CL02:2023 Medical laboratories — Requirements for quality and competence (ISO 15189:2022, IDT) and CNAS-CL02-A001:2023 Application Requirements of the Accreditation Criteria for the Quality and Competence of Medical Laboratories. It comprises the Quality Manual, procedures, work instructions and related records/forms. |  |  |  |
| Do you have a written Quality Manual available? |  |  |  | Yes. Quality Manual ODC-ZLSC is available. |
| Do you have an independent Quality Assurance department? |  |  |  | Yes. A Quality Department has been established. |
| Please provide a short description of the QA responsibilities. | 1. QMS establishment and maintenance: participate in drafting, review, revision, issuance and retrieval of QMS documents (SOPs, Quality Manual, record templates, etc.); participate regularly in management review and oversee continued effective operation of the system. 2. Training support: assist in organizing quality-related training (GLP requirements, documentation/record practices, deviation management, etc.) and oversee implementation of personnel training and authorization. 3. Manage nonconforming work, corrective actions, preventive actions and continual improvement: |  |  |  |
| Are your activities/processes covered by approved procedures (policies, SOPs or work instructions)? |  |  |  | Yes. Activities and processes are covered by approved procedures. The company has documented ODC-CX-33 Document Control Procedure and ODC-CX-34 Record Management Procedure. Document revision follows a defined process. Activities are covered by corresponding approved procedures, SOPs or work instructions, with records retained as evidence of execution. |
| How are documents such as procedures managed at your facility? | Electronic document management system Manual document management system A combination of both N/A |  |  |  |
| How are the SOPs, policies and work instructions created, maintained, and accessed by personnel? | Creation: drafted by the process-owner department based on applicable regulations and QMS requirements, industry practice and internal experience; reviewed by the department head and Quality Manager and approved by the Laboratory Director before release. Maintenance: reviewed at least annually, or promptly revised upon process change or regulatory update. Each document records revision history and the reason for change. Obsolete versions are archived and clearly identified to prevent unintended use. Access: current paper copies are signed/stamped and issued to departments; electronic copies are stored on a shared server/internal platform with role- and department-based access. Training is provided at onboarding and upon document update, with training records retained. A master list of controlled documents is maintained to ensure traceability. |  |  |  |
| How does the business assure that unintended use of obsolete documents is prevented? | For paper copies, after a new version is released the document administrator promptly retrieves all obsolete printouts from workstations and laboratories, stamps them “Obsolete”, archives them for the defined period and then destroys them. Electronic controlled documents are stored centrally on the server; after a new version is released the previous version is superseded and ordinary users cannot access or download it. Internal audits include workplace checks to confirm that no uncontrolled or obsolete documents remain at points of use. |  |  |  |
| Are your procedures maintained by a documented revision policy? |  |  |  | - Yes. A documented revision policy is followed. ODC-CX-33 Document Control Procedure defines revision and addition of documents: - To revise a document, the applicant completes a Document Application Form for review by the original reviewer (or the Quality Manager if the original reviewer is unavailable). After Quality Manager approval, the revised draft is reviewed by the original reviewer and approved/released by the designated approver. - To add a document, the applicant completes a Document Application Form; after approval the drafting–review–approval–release process is followed. The document administrator updates the Document Control List and the Document Issue and Retrieval Register. - Handwritten alteration of QMS documents is not permitted. - Changes shall be recorded on the revision page, identifying the change and the approver. - When a document is reissued, Quality retrieves and destroys the previous version. - After document review, the document administrator prepares the Annual Document Review Record; after approval by the document approver, revision is arranged. |
| Please describe the implemented Good Documentation Practices. | - ODC-CX-33 Document Control Procedure — document establishment: - Document control: controlled / uncontrolled / obsolete copies are stamped accordingly to prevent misuse. - Unique identification: documents are numbered as “company abbreviation–category–serial number”; record forms add “R + sequence”. - Version control: version letters A/B/C… and revision numbers 0/1/2…; a version increment is required after nine revisions. - Segregation of review and approval: designated reviewers and approvers sign at defined document levels. - A defined interval is maintained between the approval date and the effective date. - Periodic review: documents are reviewed before the annual internal audit and revised as needed. - Record-entry rules: blue/black ink; corrections by single-line strikethrough with signature; correction fluid is prohibited. - Archive protection: periodic archiving with protection against mould, insects, rodents and fire; electronic and paper records are regarded as equivalent. - Defined retention: different record types have defined retention periods (e.g. examination raw data 6 years; tumour molecular testing records 20 years). - Controlled destruction: destruction after expiry requires application and approval, a destruction record and confidential handling. |  |  |  |
| Provide the following documents: | ☐ Quality Manual ☐ Organizational chart of the company ☐ Floor plan ☐ Cover page from an effective SOP / procedure / work instruction ☐ List of current policies, SOPs and/or work instructions ☐ Change-management policy or SOP ☐ Templates for validation plans and reports. |  |  |  |

| FACILITIES MANAGEMENT |  |  |  |  |
| --- | --- | --- | --- | --- |
| ### Question | ### Yes | ### No | ### N/A | ### Comments / specification |
| Does the business have physical facility security (access controls and alarms, fire alarms, etc.)? |  |  |  | Yes. Per ODC-CX-10 Facility and Environmental Management Procedure, the company is equipped with safety facilities and equipment, including emergency evacuation devices and alarm systems, readily accessible emergency showers and eyewash stations, with periodic functional verification. Access to critical areas is restricted: laboratory staff enter by access card or fingerprint; other persons enter only with authorization. Smoking is prohibited throughout the premises. The safety administrator maintains fire-fighting equipment and checks its effectiveness regularly. |
| Please give an overview of the facility size and structure (one or multiple buildings, main usage, etc.). | Per ODC-CX-10 Facility and Environmental Management Procedure, the facility is purpose-designed with zoned management; areas are separated by laboratory-grade fire-resistant boards/partitions of defined height and thickness. Per ODC-AQ-34 Company Risk Assessment and Risk Control SOP, the laboratory is operated at biosafety level 2 (BSL-2). Principal use is medical laboratory testing, including throat swabs, gastric mucosa and other specimen types. Per ODC-AQ-42 Pest Control SOP, pest-control coverage includes the laboratory perimeter, open areas and laboratory rooms. |  |  |  |
| Does the business have pest control and related contracts? |  |  |  | The company has established ODC-AQ-42 Pest Control SOP: the General Administration Department prepares the pest-control plan and oversees completion of pest-control activities; Routine in-house pest-control activities are mainly performed by the General Administration Department. |
| Are facilities regularly inspected related to pest control, and is this documented? |  |  |  | Yes. Per ODC-AQ-42 Pest Control SOP, internal pest/rodent control is performed monthly. Glue-board locations are fixed, kept clean and cleared of carcasses and debris. Placement locations and quantities shall remain effective and are documented on ODC-AQ-42-R01 Pest and Rodent Control Record. |
| Is there control of temperature and humidity in the facility, and is this documented? |  |  |  | Per ODC-CX-10 Facility and Environmental Management Procedure, testing areas with temperature or humidity requirements are equipped with thermometers/hygrometers. On-duty laboratory staff record temperature/humidity daily. Freezers/refrigerators used for reagents or samples, incubators and similar equipment have thermometers or probes. A designated person assigned by the Laboratory Director records temperatures and checks the monitoring system. Records are kept on ODC-CX-10-R01 Indoor Temperature and Humidity Record. |
| What actions are taken if temperature/humidity is out of specification? | Per ODC-CX-10 Facility and Environmental Management Procedure: When temperature or humidity is out of specification, prompt adjustment is required. Temperature is adjusted via HVAC; a dehumidifier is used if humidity is high and a humidifier if humidity is low. If specification cannot be restored, the departmental equipment administrator is notified to arrange maintenance; testing is stopped during repair and the Nonconformity Identification and Control Procedure is executed. For reagent/sample storage units, incubators and similar equipment, if specification cannot be restored, stored materials are transferred to other qualified units during repair. Upon detecting an excursion, the Laboratory Director is notified immediately; the quality supervisor completes the Facility and Environment Out-of-Control and Handling Record. If result quality may be affected, testing is stopped and the Nonconformity Identification and Control Procedure is applied. |  |  |  |
| Does the business have back-up power generators? |  |  |  | No dedicated back-up generator is installed. Dual-feed electrical supply is available (manual switchover, not automatic). Five UPS units are connected to critical equipment. |
| If yes, are they tested on a regular basis and are the results documented? |  |  |  | Per the ODC-AQ Electrical Safety Standard, all electrical installations, equipment and appliances shall be periodically maintained by qualified electrical personnel. UPS-related records include: ODC-AQ-27-R01 UPS Maintenance Record; ODC-AQ-27-R02 UPS Maintenance Plan. |
| Does your company have a business continuity / disaster recovery plan? |  |  |  | Per ODC-CX-10 Facility and Environmental Management Procedure, emergency measures are defined for unexpected water or power interruption. A laboratory safety risk assessment is performed annually and is re-performed when facilities change. Per ODC-CX-32 Emergency Preparedness Procedure, the General Administration Department prepares the internal Emergency Drill Plan each January for approval by the Laboratory Director and subsequent implementation. |
| If yes, is it tested on a regular basis and are the results documented? | Yes, all are documented. - ODC-CX-32-R01 Emergency Drill Plan; - ODC-CX-32-R02 Emergency Drill Record — the drill organizer plans materials, participants, timing, process and recording method; after the drill the form is completed, reviewed by the safety responsible person and archived. |  |  |  |
| Has a risk-based approach been taken related to the business continuity / disaster recovery plan(s), and is this documented? |  |  |  | The laboratory has established ODC-CX-08 Risk Management Procedure and periodically performs identification and assessment of business-interruption risks. Business continuity and disaster-recovery strategies are derived from that risk assessment. Related assessments, plans, drills and reviews are documented; plans are reviewed and updated periodically as risks change. |

| TRAINING |  |  |  |  |
| --- | --- | --- | --- | --- |
| Question | Yes | No | N/A | Comments / specification |
| Are there up-to-date job descriptions and CVs available for all personnel? |  |  |  | - Yes. Per ODC-CX-09 Human Resource Management Procedure, the Administrative and HR Department integrates human resources, maintains and timely updates personnel files for all staff, and makes them readily accessible for consultation. |
| Which of the following trainings are provided to the staff? | GLP training GCP / GCLP training GDPR / Data Protection Act training — we have a confidentiality management procedure GMP training HIPAA training Good Documentation Practice training |  |  |  |
| Does your company have a general training process in place? |  |  |  | Yes. Per ODC-CX-09 Human Resource Management Procedure, a documented general training process is in place, forming a closed loop from training-plan development, implementation, record archiving and assessment through monitoring and retraining. |
| Please explain the overall training process (e.g. instructor-led training, on-the-job training, e-learning; self-training, external training). | Per ODC-CX-09 Human Resource Management Procedure, training is delivered via lectures, reading/self-study, other methods (e.g. online meetings) and external training. |  |  |  |
| Does your company have a function-specific training procedure in place? |  |  |  | Yes. ODC-CX-09 Human Resource Management Procedure has been established. |
| Please explain the function-specific training process (e.g. instructor-led training, on-the-job training, e-learning; self-training, external training). | Per ODC-CX-09 Human Resource Management Procedure, training is delivered via lectures, reading/self-study, other methods (e.g. online meetings) and external training. Training process: the laboratory training plan is compiled, checked, adjusted, approved and posted by the document administrator; training notices are issued in advance; the administrator manages attendance, collection/archiving of materials and updates to personnel training histories. Trainees shall arrange work accordingly and report in advance if unable to attend. After training, the trainer conducts assessment and submits training records. The safety responsible person periodically organizes safety training and emergency drills. Unplanned internal/external training requires stepwise approval; after external training, effectiveness evaluation is performed and materials are archived; absentees shall receive make-up training and assessment. |  |  |  |
| Is employee training documented, and are these records maintained? |  |  |  | Per ODC-CX-09 Human Resource Management Procedure: 1) After each training session, the document administrator verifies completeness of the Attendance Sheet, Training Record and examination papers, files them by category and registers them in the Employee Training and Continuing Education History. 2) After external training/meetings, effectiveness is evaluated by Quality together with the technical group leader against the External Training/Meeting Approval Form; after archiving, the Employee Training and Continuing Education History is updated. 3) Records are retained with personnel files for 25 years per ODC-CX-34 Record Management Procedure; hard-copy originals are archived to the records room annually. |
| How are these records retained? | Records are retained with personnel files for 25 years per ODC-CX-34 Record Management Procedure; hard-copy originals are archived to the records room annually. |  |  |  |
| How long are training records retained? | Training records are retained for 25 years. |  |  |  |
| Is employee training conducted on a continuing basis? |  |  |  | Per ODC-CX-09 Human Resource Management Procedure, the Quality Department convenes relevant departmental staff for training monthly. |
| How often do employees receive retraining, and how is it initiated? | New employees receive at least two competency assessments within the first 6 months, with records. Retraining and reassessment are required upon change of duties, return to work after more than 6 months’ absence, or changes to policies, procedures or technology. Personnel who do not pass assessment shall be retrained and reassessed and may resume duty only after qualification, with records. |  |  |  |
| If contract personnel (interim staff, consultants) are employed, are training requirements equivalent to those for full-time employees? |  |  |  | No contract personnel (interim staff or consultants) are employed. |
| Will CVs, job classifications and training records be accessible upon request? |  |  |  | Per ODC-CX-09 Human Resource Management Procedure, the Administrative and HR Department maintains personnel files (including training records) of all active staff and makes them readily accessible. Therefore CVs, job classifications and training records can be provided upon request. |

| AUDITING AND VENDOR MANAGEMENT |  |  |  |  |
| --- | --- | --- | --- | --- |
| Question | Yes | No | N/A | Comments / specification |
| Is a procedure in place for initial vendor/supplier qualification? |  |  |  | Per ODC-CX-16 External Service and Supply Management Procedure, after investigation, evaluation and review, suppliers are added to the Qualified Service Provider and Supplier List, with qualification records archived. |
| Is a procedure in place for reassessment/requalification of vendors/suppliers? |  |  |  | Per ODC-CX-16 External Service and Supply Management Procedure, annual supplier evaluation is organized at the beginning of each year, with attention to qualification validity and changes in capability. Corrective action may be required if issues are found; qualification is withdrawn if requirements are still not met. |
| Are vendors/suppliers reassessed following a risk-based approach? |  |  |  | Reassessment of suppliers and contracted laboratories is performed according to their impact on examination quality and actual performance, with different frequencies and triggers, as defined in ODC-CX-16 External Service and Supply Management Procedure. |
| Are vendor/supplier audits performed? |  |  |  | Per ODC-CX-16 External Service and Supply Management Procedure, supplier review is organized at least annually and ODC-CX-16-R03 Supplier Investigation and Review Form is completed. |
| How frequently are vendors/suppliers evaluated? | ODC-CX-16 External Service and Supply Management Procedure specifies that periodic evaluation is conducted once per year. |  |  |  |
| Are internal audits performed? |  |  |  | A complete internal audit process is defined in ODC-CX-39 Internal Audit Procedure. |
| Are internal audits planned and performed following a risk-based approach? |  |  |  | Per ODC-CX-39 Internal Audit Procedure, audit intensity and frequency are increased for departments that are more critical or that have had more findings, based on departmental importance and previous audit results, with focus on areas of key significance to service users (sections 4.3.5 and 4.6.1), reflecting a risk-based planning and execution approach. |
| Is an external audit program in place? |  |  |  | No external audit program is established. |

| 6. INCIDENT, CAPA AND COMPLAINT MANAGEMENT |  |  |  |  |
| --- | --- | --- | --- | --- |
| Question | Yes | No | N/A | Comments / specification |
| Do you have an incident and deviation management system in place? |  |  |  | Yes. ODC-CX-29 Nonconformity Identification and Control Procedure and ODC-CX-08 Risk Management Procedure have been established, covering identification, recording and handling of incidents and deviations. |
| What is your follow-up process for incidents and deviations? | Per ODC-CX-37 Corrective Action Control Procedure, the process is: identification → root-cause investigation → development of corrective actions → approval → implementation → effectiveness verification → record archiving → input to management review. |  |  |  |
| Do you have a (external) complaint management system in place? |  |  |  | Yes. ODC-CX-31 Complaint Management Procedure has been established, covering receipt, handling, response and follow-up verification of internal and external complaints. |
| What is your follow-up process for complaints? | Per ODC-CX-31 Complaint Management Procedure: receive and record → assign responsible department → investigate and analyse → classify valid/invalid complaints → develop corrective actions → review and approve → respond to the customer → follow-up verification → statistical analysis → identify improvement opportunities. |  |  |  |
| Are documented procedures in place for incident, deviation and complaint handling? |  |  |  | Yes. Written procedures include ODC-CX-37 Corrective Action Control Procedure, ODC-CX-31 Complaint Management Procedure, ODC-CX-29 Nonconformity Identification and Control Procedure, etc. |
| Are risks assessed for incident, deviation and complaint handling? |  |  |  | ODC-CX-31 Complaint Management Procedure requires analysis of whether risk factors exist. In ODC-CX-08 Risk Management Procedure, “nonconformities and corrective actions” and “complaints” are listed as risk sources in the risk identification register. |
| Are CAPA (corrective actions and preventive actions) recorded and followed up? |  |  |  | ODC-CX-37 Corrective Action Control Procedure requires completion of the Nonconformity and Corrective Action Record, which is archived after approval, implementation and effectiveness verification. |
| Does your company provide investigation reports? |  |  |  | Investigations are initiated per ODC-CX-37 Corrective Action Control Procedure, producing the Nonconformity and Corrective Action Record (including root cause, impact assessment, CAPA and measures to prevent recurrence). Complaints are recorded on the Complaint Handling Form per ODC-CX-31 Complaint Management Procedure. |
| How are CAPA related to investigation reports communicated? What are the timelines for providing a final investigation report? | Per ODC-CX-31 Complaint Management Procedure, the responsible department completes the Complaint Handling Form and submits root-cause analysis and corrective actions to the quality/technical responsible person for review; after approval the form is sent to Quality, Customer Service replies to the customer, and Quality tracks verification of actions. In principle all complaints are closed within three working days; complex complaints may be extended. General complaints are acknowledged within 4 hours and major complaints within 1 hour. Per ODC-CX-37 Corrective Action Control Procedure, the responsible department/technical group develops corrective actions. Routine departmental measures are approved by the department head; company-level general measures by the Quality Responsible Person; measures involving major resource input or major impact by the General Manager. Cross-departmental actions are coordinated and monitored by the Quality Responsible Person. The Nonconformity and Corrective Action Record is completed within the predefined timeframe. Per ODC-CX-39 Internal Audit Procedure, internal auditors issue the Nonconformity and Corrective Action Record to the responsible department; the department analyses the cause, proposes corrective actions and a timeline, and implements them after approval by the Quality Responsible Person. Internal auditors verify effectiveness; incomplete actions are reported to the audit team leader, and actions that cannot be completed are escalated to management. Internal audit reports are written by the Quality Manager and released after approval by the Laboratory Director. |  |  |  |

| 7. DOCUMENT AND DATA INTEGRITY / RETENTION AND ARCHIVING |  |  |  |  |
| --- | --- | --- | --- | --- |
| Question | Yes | No | N/A | Comments / specification |
| How is record keeping performed? | Electronic Paper-based Electronic and paper-based |  |  |  |
| Do you have a record retention plan/procedure? |  |  |  | Yes. ODC-CX-34 Record Management Procedure defines full lifecycle management of records: identification, entry, archiving, retention, retrieval and destruction. |
| Please provide the retention period for your documentation. If this differs by document type, please provide details. | Retention periods differ by record type, as shown in the table below: Record type; Retention period. All pharmaceutical clinical trial related records; 25 years (or per customer requirement). |
| Are records stored in a manner that will protect their integrity for the required time period? |  |  |  | Per ODC-CX-34 Record Management Procedure, records shall be stored in a suitable environment (protection against mould, insects, rodents and fire). Electronic records are maintained and backed up per the Computer Hardware and Software Management Procedure. Temporary files are deleted promptly to prevent loss, damage or unauthorized modification. |
| Do you have a management system in place for archiving? |  |  |  | A complete archive management system is established under ODC-CX-34 Record Management Procedure. Records of the previous year are archived in the first quarter; the Record Archiving Registration Form is completed and the Archived Records List is updated. Technical groups collate and bind records by category; the archive administrator signs for receipt. |
| Are records retained in the facility or stored at an external archiving facility? |  |  |  | Retained in-house. Per ODC-CX-34 Record Management Procedure, records are held by departments and technical groups, temporarily stored locally, and periodically transferred to laboratory-managed archives. |
| Are records stored at an internal archive? |  |  |  | Per ODC-CX-34 Record Management Procedure, records are archived in the laboratory’s internal archive room and managed uniformly by the document administrator or information administrator. |
| Are records stored at an external archiving facility? |  |  |  | Records are not stored at an external archiving facility. |
| Has the external archiving facility been audited by a company representative? |  |  |  | Not applicable. |
| Does your business manage confidentiality and privacy for subject-related data? |  |  |  | Yes. ODC-CX-04 Confidentiality Management Procedure has been established, specifying protection of patients’ personal information, examination results, examination data and other confidential information. All personnel shall sign a Confidentiality Agreement. |
| How is confidentiality and privacy assured? | Per ODC-CX-04 Confidentiality Management Procedure, Protection measures include: - All personnel (internal and external) sign a Confidentiality Agreement; - Information is classified and access is authorized; unauthorized persons shall not have access; - Computer information is encrypted or access-controlled and shall not be connected to networks without approval; - Unrelated persons shall not enter examination areas; visitors are registered on entry and exit; - Customers shall not observe unrelated testing or take photographs/videos; - The General Manager or Quality Responsible Person inspects confidentiality implementation quarterly. |  |  |  |
| Are data changes documented in the systems used (i.e. audit trail), including an editing trail with date/time stamping? |  |  |  | A LIS is not used for this project. The study-related ABI 7500 PCR instrument and ABI 3500DX sequencer both have audit-trail functionality. |
| Does the business have computer-generated, time-stamped audit trails to track creation, modification and deletion of records that describe when, by whom and why changes were made, without obscuring original information? |  |  |  | Not applicable to this project; a LIS is not used. |
| How long can data be electronically stored in and retrieved from the system? | Not applicable to this project; a LIS is not used. |  |  |  |

| 8. CHANGE MANAGEMENT |  |  |  |  |
| --- | --- | --- | --- | --- |
| ### Question | ### Yes | ### No | ### N/A | ### Comments / specification |
| Is there a procedure in place for change management? |  |  |  | Change management is implemented across multiple procedures, covering documents (ODC-CX-33 Document Control Procedure), examination procedures (ODC-CX-20 Selection, Verification and Validation of Examination Procedures), IT systems (ODC-CX-30 Computer Hardware and Software Management Procedure), personnel (ODC-CX-09 Human Resource Management Procedure) and reagent suppliers (ODC-CX-14 Reagent and Consumable Management Procedure). |
| Is a risk-based approach implemented during the change management process? |  |  |  | Per ODC-CX-09 Human Resource Management Procedure, personnel/role changes require revision and approval of job descriptions; accreditation-related changes are notified to CNAS within 20 working days; retraining and reassessment are required upon role change, return after more than 6 months’ absence, or policy/procedure/technology change. The procedure does not explicitly use the wording “risk-based approach”. Per ODC-CX-10 Facility and Environmental Management Procedure, a laboratory safety risk assessment is performed annually and is re-performed when facilities change, with reference to the Risk Management Procedure. Change-triggered re-assessment reflects a risk-based approach. Per ODC-CX-12 Equipment Management Procedure, application and approval are required before return to use after repair, storage or transfer; after fault repair, calibration, IQC, comparison or retained-sample retest is selected according to the impact on analytical performance. Status labelling is qualified / restricted use / out of service. The procedure does not explicitly use the wording “risk-based approach”. Per ODC-CX-14 Reagent and Consumable Management Procedure, supplier changes require re-evaluation; reagents/consumables of different lots/shipments shall be verified before use; nonconforming materials are returned and used as input to supplier evaluation; quality risk of materials within 2 months of expiry is assessed before return/exchange. Per ODC-CX-20 Selection, Verification and Validation of Examination Procedures, examination-procedure changes are approved on an Examination Procedure Change Approval Form; whether existing verification/validation protocols can be used is assessed; reagent upgrades, instrument updates and calibrator-traceability changes are verified as a new system; partial performance verification is required after events that severely affect analytical performance. The scope of verification is determined by the impact of the change, reflecting a risk-based approach. Per ODC-CX-33 Document Control Procedure, document revision and addition follow the Document Application Form for review by the original reviewer and approval/release by the approver. Risk grading is not defined, nor are risk-based review frequency or methods specified. |
| Is the change management process implemented throughout the company, relating to all activities (e.g. project changes, IT changes)? | The company implements unified change control. Change management is defined in multiple procedures covering documents (ODC-CX-33 Document Control Procedure), examination procedures (ODC-CX-20 Selection, Verification and Validation of Examination Procedures), IT systems (ODC-CX-30 Computer Hardware and Software Management Procedure), personnel (ODC-CX-09 Human Resource Management Procedure) and reagent suppliers (ODC-CX-14 Reagent and Consumable Management Procedure). |  |  |  |
| Are all changes approved before execution? |  |  |  | Yes. All changes shall be approved by the designated approver before execution (per ODC-CX-33 Document Control Procedure; ODC-CX-20 Selection, Verification and Validation of Examination Procedures; ODC-CX-30 Computer Hardware and Software Management Procedure). |
| Are all changes managed in a system? |  |  |  | Per ODC-CX-30 Computer Hardware and Software Management Procedure, software go-live/authorization uses the Computer Software Validation Record; new/changed tests use the System New Examination Item Application Form and System Examination Item Information Change Record; access-right changes use the Information System Access Rights Record; database changes use the Database Server Maintenance, Upgrade and Change Record. However, UPS maintenance, breakdown repair and workstation rounds still use paper forms and are not fully systematized. Per ODC-CX-09 Human Resource Management Procedure, recruitment applications and personnel changes are processed via Administrative and HR; training plans, competency assessments and personnel files are paper forms. Per ODC-CX-10 Facility and Environmental Management Procedure, facility/environment out-of-control records, temperature/humidity records, water-quality monitoring, housekeeping checks and visitor registers are paper records. Per ODC-CX-12 Equipment Management Procedure, equipment procurement follows OA application under the External Service and Supply Management Procedure; equipment status labelling, acceptance, repair, calibration and decommissioning/scrapping use paper forms. Per ODC-CX-14 Reagent and Consumable Management Procedure, purchase requests go through the OA system; inbound, outbound and inventory use the “reagent and consumable intelligent management system”; paper forms are used only when the system is unavailable. Acceptance, QC, scrapping and returns still use paper forms. Per ODC-CX-20 Selection, Verification and Validation of Examination Procedures, examination-procedure change requests are raised in OA; reagent purchases for new projects are requested in OA. However, the Examination Procedure Change Approval Form, New Technology and New Project Initiation Review Form and performance verification reports remain paper/file-based. Per ODC-CX-33 Document Control Procedure, document revision, addition, review, issue, retrieval, destruction and loan all use paper forms (Document Application Form, Document Change Register, Document Destruction Application and Record, etc.); electronic copies are stored only on designated computers. |
| Is an effectiveness review performed after finalization of changes? |  |  |  | Per ODC-CX-09 Human Resource Management Procedure, upon role change, return after more than 6 months’ absence, or policy/procedure/technology change, personnel shall be retrained and reassessed and may resume duty only after qualification, with records. Per ODC-CX-10 Facility and Environmental Management Procedure, a laboratory safety risk assessment is performed annually and is re-performed when facilities change (4.5); environmental out-of-control handling shall be recorded and followed up. Per ODC-CX-12 Equipment Management Procedure, after repair, return from storage or transfer, the technical group leader shall apply; after review by the Technical Director and approval by the Laboratory Director the equipment may be placed in service. After fault repair, the cause is analysed and calibration, IQC, comparison or retained-sample retest is performed according to impact. Per ODC-CX-14 Reagent and Consumable Management Procedure, when the supplier changes, the new supplier shall pass evaluation before being added to the supplier list; reagents/consumables of different lots/shipments shall be verified before use; nonconforming materials are returned and used as input to supplier evaluation. Per ODC-CX-20 Selection, Verification and Validation of Examination Procedures, examination-procedure changes require an Examination Procedure Change Approval Form, reviewed by the Technical Director or Quality Responsible Person and approved by the Laboratory Director. The technical group leader assesses whether existing verification/validation protocols can be used; re-verification or validation is required after the change. Per ODC-CX-30 Computer Hardware and Software Management Procedure, software is tested and confirmed in writing by the Laboratory Director before go-live; after new/changed tests, the applicant and Laboratory Director sign confirmation; access-right changes are signed, enabled in the system and archived. Per ODC-CX-33 Document Control Procedure, document revisions are reviewed by the original reviewer and approved/released by the designated approver. After review results are confirmed, the document administrator prepares the Annual Document Review Record and arranges revision after approval. |
| Is an effectiveness review performed after finalization of changes? |  |  |  | ODC-CX-30 Computer Hardware and Software Management Procedure: software is tested and confirmed by the Laboratory Director before go-live; after new/changed tests the applicant and Laboratory Director sign confirmation; access-right changes are signed, enabled and archived. ODC-CX-09 Human Resource Management Procedure: upon role change, return after more than 6 months’ absence, or policy/procedure/technology change, personnel shall be retrained and reassessed and may resume duty only after qualification. ODC-CX-10 Facility and Environmental Management Procedure: a laboratory safety risk assessment is performed annually and is re-performed when facilities change, with reference to the Risk Management Procedure. ODC-CX-12 Equipment Management Procedure: after repair, return from storage or transfer, equipment may be used only after application by the technical group, review by the Technical Director and approval by the Laboratory Director; after fault repair, calibration, IQC, comparison or retained-sample retest is performed according to impact. ODC-CX-14 Reagent and Consumable Management Procedure: when the supplier changes, the new supplier shall pass evaluation before being added to the supplier list; reagents/consumables of different lots/shipments shall be verified before use. ODC-CX-20 Selection, Verification and Validation of Examination Procedures: examination-procedure changes require an Examination Procedure Change Approval Form, reviewed by the Technical Director or Quality Responsible Person and approved by the Laboratory Director; re-verification or validation is required after the change; examination-procedure performance is reviewed annually. ODC-CX-33 Document Control Procedure: document revisions are reviewed by the original reviewer and approved/released by the designated approver; after review results are confirmed, the document administrator prepares the Annual Document Review Record and arranges revision after approval. |

| 9. SYSTEM VALIDATION |  |  |  |  |
| --- | --- | --- | --- | --- |
| ### Question | ### Yes | ### No | ### N/A | ### Comments / specification |
| Does the business conduct validation of systems to ensure accuracy, reliability, consistent intended performance, and the ability to discern invalid or altered records? |  |  |  | Per ODC-CX-12 Equipment Management Procedure, engineers install, commission, verify and calibrate instruments. After acceptance, the manufacturer’s engineer completes an installation report and issues a verification/calibration certificate. Paper records use strikethrough corrections with signature/annotation to prevent tampering and remain traceable for 25 years. |
| Is there a procedure for validation and/or qualification of equipment/devices? |  |  |  | Per ODC-CX-12 Equipment Management Procedure, newly purchased equipment shall complete acceptance and performance verification before use; after repair, equipment shall be verified as satisfactory before return to service. |
| Do you have an equipment-failure back-up plan? |  |  |  | Per ODC-CX-32 Emergency Preparedness Procedure, in the event of instrument failure the supplier provides an equivalent-model replacement as an emergency back-up; the replacement shall complete performance verification before use. |
| Is there a procedure for validation and/or qualification of computerized systems? |  |  |  | Yes. ODC-CX-30 Computer Hardware and Software Management Procedure has been established, defining software validation requirements and record forms. |
| What type of risk-based approach is in place for validation of computer systems? | Per ODC-CX-30 Computer Hardware and Software Management Procedure, validation addresses functionality, performance, usability, compatibility and security. A risk-based approach requires risk classification according to impact on examination quality, patient safety and data integrity, and determines the scope and depth of validation accordingly. |  |  |  |
| Are the computerized systems validated to ensure data integrity? |  |  |  | Per ODC-CX-30 Computer Hardware and Software Management Procedure, software validation shall ensure protection of sensitive data and prevent unauthorized access; system additions or changes shall be verified; database backups shall be verified for integrity. A risk-based approach is used, focusing on data acquisition, transmission, storage, modification, audit trail, backup/restore and access control to ensure data integrity. |
| Are computerized systems validated? - In compliance with GCP / GCLP standards? - In accordance with 21 CFR Part 11? - In accordance with GAMP 5 validation? |  |  |  | N/A, unless the relevant system is used for FDA-regulated electronic records/electronic signatures. Where applicable, assessment can be performed according to project requirements. A risk-based computerized-system validation approach is used, referring to GAMP 5 risk-management and validation-lifecycle concepts. |
| Please provide a list of compliant computerized systems (addendum). |  |  |  |  |

| 10. LABORATORY TEST SPECIFIC QUESTIONS |  |  |  |  |
| --- | --- | --- | --- | --- |
| ### Question | ### Yes | ### No | ### N/A | ### Comments / specification |
| ### Samples |  |  |  |  |
| Is there a procedure in place for sample reception? |  |  |  | Yes. ODC-CX-19 Primary Sample Reception and Handling Procedure has been established, defining sample reconciliation, receipt, information entry, rejection of unsuitable samples and testing under concession. |
| Can you guarantee complete cold-chain custody? |  |  |  | Per ODC-CX-18 Primary Sample Collection and Transport Procedure, samples are transported in rigid specimen boxes under temperature control, with transport temperature recorded. Receipt includes a transport-temperature check; if temperature is out of control, impact on results is assessed. For storage, per ODC-CX-32 Emergency Preparedness Procedure, a remote real-time temperature/humidity monitoring system is in place and abnormalities are handled immediately. |
| Do your processes describe sample integrity checks? |  |  |  | Per ODC-CX-19 Primary Sample Reception and Handling Procedure, incoming samples are checked against acceptance criteria, including integrity, identification and collection time. Unsuitable samples are rejected or, after assessment, tested under concession with a documented remark in the system. |
| Do you have a sample identification process (accession labels, barcode, readers, re-labelling) in place? |  |  |  | Samples are identified by a unique 12-digit barcode, ensuring traceability from collection, receipt, transfer, distribution and examination through reporting (ODC-CX-19 Primary Sample Reception and Handling Procedure). Pathology reports use the pathology number as the unique identifier (ODC-CX-26 Examination Result Reporting Procedure). |
| Does your company have a procedure for sample storage and retention? |  |  |  | How long are samples retained following analysis? Per ODC-SOP-CRO-01-012 Clinical Research Group Sample Reception and Handling Procedure, post-analysis sample retention follows the clinical trial protocol, ethics approval and customer requirements; samples are retained long-term until project completion. |
| Do you have a process in place for sample disposal? |  |  |  | Per ODC-SOP-CRO-01-019 Clinical Research Group Post-Examination Sample Management Procedure, samples exceeding the retention period are disposed of safely as medical waste by the respective technical groups. |
| Is there a temperature monitoring system? |  |  |  | Per ODC-CX-32 Emergency Preparedness Procedure, laboratory storage equipment is equipped with remote real-time temperature/humidity monitoring. Per ODC-CX-10 Facility and Environmental Management Procedure, each technical group performs routine temperature/humidity monitoring and recording according to instrument and reagent requirements; out-of-control conditions are handled immediately. |
| What storage temperatures are offered? | Room temperature 2 – 8 °C -20 °C -80 °C ☐ Vapour-phase liquid nitrogen |  |  |  |
| Is there a system for continuous temperature monitoring of refrigerators / freezers / rooms, etc.? If yes, is this temperature monitoring system connected to a 24/7 alarm system? |  |  |  | Per ODC-CX-32 Emergency Preparedness Procedure, laboratory storage equipment has remote real-time temperature/humidity monitoring. When equipment failure causes temperature to reach the alarm limit, the system simultaneously sends WeChat alerts to multiple persons, providing 24-hour monitoring and alarm notification. |
| Is there a spill procedure in place? |  |  |  | Per ODC-CX-10 Facility and Environmental Management Procedure, in the event of contaminant leakage or spillage, all contaminated surfaces (benches, floors, walls, etc.) and air are disinfected according to procedure. Chemical spills are handled per the emergency process in the Safety Manual (ODC-CX-32 Emergency Preparedness Procedure). |
| ### Equipment used and equipment validations |  |  |  |  |
| Is an equipment database maintained? |  |  |  | Per ODC-CX-12 Equipment Management Procedure, the laboratory shall maintain an Equipment List and Equipment Basic Information Register covering identification, manufacturer information, calibration records and maintenance/repair records. Files are managed by the technical groups or the General Administration Department. |
| Please add a copy of the list of equipment. |  |  |  |  |
| Are all laboratory instruments validated prior to use? |  |  |  | Per ODC-CX-12 Equipment Management Procedure, newly purchased equipment shall complete acceptance and performance verification, and equipment shall be verified after repair before reuse. Per ODC-CX-20 Selection, Verification and Validation of Examination Procedures, all examination procedures undergo independent verification before routine use. |
| Are all laboratory instruments 21 CFR Part 11 compliant, or have they been evaluated against 21 CFR Part 11? |  |  |  | Study-related instruments including the ABI 7500 PCR system, ABI 3500DX sequencer and microplate reader are 21 CFR Part 11 compliant. |
| Do you have a process in place in case of instrument switch mid-trial? |  |  |  | Per ODC-CX-32 Emergency Preparedness Procedure, in the event of instrument failure the supplier provides an equivalent-model replacement as an emergency back-up; the replacement shall complete performance verification before use. |
| ### Method validations |  |  |  |  |
| Are there procedures in place for method validation of test(s)? |  |  |  | Yes. ODC-CX-20 Selection, Verification and Validation of Examination Procedures has been established, defining requirements for method verification and validation. |
| Please describe your method validation process. | Per ODC-CX-20 Selection, Verification and Validation of Examination Procedures, the verification/validation process is: prepare a protocol (including performance characteristics and acceptance criteria) → Technical Director approval → implementation by authorized personnel with raw-data retention → draft verification/validation report → technical group leader review → Technical Director approval. |  |  |  |
| Is a validation plan and report available for each test(s)/parameter(s) performed by your laboratory? |  |  |  | Per ODC-CX-20 Selection, Verification and Validation of Examination Procedures, each new examination requires an Analytical Performance Verification Report or Analytical Performance Validation Report, reviewed and approved before implementation. Performance of approved examination procedures is reviewed at least annually. |
| Please explain, with rationale, when verification/validation is performed and when it is not. | Per ODC-CX-20 Selection, Verification and Validation of Examination Procedures, verification is performed: before routine use; after events that severely affect analytical performance (e.g. major instrument component failure, relocation, severe environmental loss of control); and when elements of the measuring system change (reagent upgrade, instrument update, change in calibrator traceability, etc.). Validation is required for: non-standard methods; laboratory-developed methods; standard methods used outside their intended scope; and modified validated methods. |  |  |  |
| Upon request, would you be able to provide the validation report(s) for test(s)/parameter(s)? |  |  |  | Yes. Per ODC-CX-20 Selection, Verification and Validation of Examination Procedures, verification/validation reports are prepared by the test owner, reviewed by the technical group leader, approved by the Technical Director and archived by the document administrator. Reports can be provided upon request. |
| Are there policies stating that validation reports cannot be made available to clients / the Trial Master File (TMF)? |  |  |  | To be answered by VCB. |
| ### Analytical procedures |  |  |  |  |
| Are your SOPs for all test(s)/parameter(s) compliant with GCP / ISO 15189 / CAP-CLIA requirements? |  |  |  | Compliant with ISO 15189 requirements. |
| Would you provide the method, instrument, specimen requirement and volume for test(s)/parameter(s)? |  |  |  | Specific requirements are defined in ODC-XMSC Project Manual and the corresponding work instructions (per ODC-CX-18 Primary Sample Collection and Transport Procedure and ODC-CX-17 Examination Request Management Procedure). |
| Is a procedure in place for testing revalidation for the test(s)/parameters? |  |  |  | Per the relevant procedures, verification or adjustment is required when: the examination procedure changes (re-verification or validation, ODC-CX-20); the reagent lot changes (verification, ODC-CX-14 Reagent and Consumable Management Procedure); or the QC lot changes (re-accumulation of mean and SD, ODC-CX-23 Internal Quality Control Procedure). |
| If Yes, how often does this occur? | Examination-procedure change: as needed (event-driven) - Reagent lot change: at each lot change - QC lot change: at each lot change (re-evaluation after accumulating 20 days of data) |  |  |  |
| Are you participating in external quality programs for test(s) or parameter(s) performed by your laboratory? |  |  |  | Per ODC-CX-24 Proficiency Testing and Interlaboratory Comparison Procedure, the laboratory preferentially participates in PT/EQA schemes organized by the NCCL (National Center for Clinical Laboratories) and the Shanghai Center for Clinical Laboratory. |
| Please explain, with rationale, when External Quality Assessment (EQA) samples are run and when they are not. | Per ODC-CX-24 Proficiency Testing and Interlaboratory Comparison Procedure, EQA/PT applies to examinations with an available EQA/PT scheme and to newly introduced examinations (new tests preferentially use measurement audit or interlaboratory comparison). Where no PT scheme exists and interlaboratory comparison is not suitable, alternative approaches may be used after authorization and shall be documented. |  |  |  |
| Do you run internal quality controls (IQCs) for test(s)/parameter(s) performed by your laboratory? |  |  |  | Per ODC-CX-23 Internal Quality Control Procedure, in principle IQC is performed for all examinations offered; at least one QC determination is run on each testing day or each analytical run. |
| Do you run internal quality controls (IQCs) for test(s)/parameter(s) performed by your laboratory? | Per ODC-CX-23 Internal Quality Control Procedure, in principle IQC is performed for all examinations offered; at least one QC determination is run on each testing day or each analytical run. |  |  |  |
| Is any analytical instrument directly interfaced with your LIMS? |  |  |  | Analytical instruments are not directly interfaced with a LIMS. |
| In what format is the data transferred to Cerba Research / other vendors? | Data are transferred via the Datalink website in PDF and JPG formats. |  |  |  |
| Is there a periodic assessment of the competency of laboratory employees for the analytical procedures? (Should be Yes; consider adding VCB assessment of technicians) |  |  |  | Per ODC-CX-09 Human Resource Management Procedure, periodic assessment of laboratory personnel competency for analytical procedures is required: - Existing staff: at least one competency assessment per year. - New staff: at least two competency assessments within the first 6 months. - Assessment covers technical skills and theoretical knowledge; staff who do not pass shall be retrained and reassessed before working independently. - Before this project starts, the customer provides study-specific training and competency assessment covering the project method, sample receipt/handling, raw-data recording, compliance requirements and deviation handling. Personnel who pass theoretical and practical assessment are authorized by the customer. Only authorized personnel may perform project testing; unauthorized personnel shall not perform experiments independently. If methods or project requirements subsequently change, the customer may arrange supplementary training and reassessment to ensure continued competence. |
| Do you hold any other scientific expertise concerning the requested test(s)? |  |  |  | The laboratory has end-to-end technical capability for selection, verification and validation of examination procedures, internal QC, proficiency testing and internal comparison (see ODC-CX-20, ODC-CX-23, ODC-CX-24, ODC-CX-25, etc.). Technical groups include molecular, pathology, microbiology and clinical research (ODC-CX-33 Document Control Procedure). |
| Are you working with other subcontractors for specific tests provided to Cerba Research? |  |  |  | NA |
| If yes, identify the tests for which subcontractors will be used and name the subcontractors. Cerba Research will also need to qualify these subcontractors. | Test |  |  | Subcontractor |
|  | N/A |  |  | N/A |
|  | N/A |  |  | N/A |
|  | N/A |  |  | N/A |
|  | N/A |  |  | N/A |